## Slide 1
Time to MA vs A10 Analysis – Heparin Pivotal Study
February – March 2024
2024

## Slide 2: Overview
Objective: Assess the distributions of the Time to MA (post R time) relative to 10 min (time to A10 parameter)
Data
  HN pivotal trial
    HN cartridge
      CFFH and CRTH assays only
Primary and Secondary Analysis
  Primary Analysis: Entire Population
  Subgroups:
    Study Indication Subgroups: Liver transplantation surgery, On-pump surgery, Not on-pump surgery
    Study Timepoint Subgroups: 001,002,003
    Coagulopathy Subgroups: Normal, Hypo, Hyper – according to the site physician assessments

## Slide 3: Analysis of Time to MA
Definitions:
Endpoint:
Reference point: 10 min (A10 time)
Statistical estimate
Distributions of : mean, SD, P > 10
Conclusion (jumping ahead):
CFFH MA, mean = 12.11, SD = 3.71
CRTH MA, mean = 19.78, SD = 3.63
Mean of CRTH is significantly larger than that of 10 minutes. With probability of ≈ 99% ,
is greater then 10 minutes.
Mean of CFFH is larger than 10 minutes. With probability of ≈ 72% ,
is greater then 10 minutes.

## Slide 4
Plot 1
  Two approaches of estimating distributions
  Red: Model based distribution (mixed effect model)
  Blue: Data summary
  Summary Statistics
  Red dashed line: a threshold of 10 minutes
  Shaded sky-blue area: the probability of observing Time to MA greater than 10 minutes
  Vertical grey dashed line: Average and 95 % CI of

## Slide 5
Plot 2
  Brown line - the fitted Mixed Effects model
  Red dashed line - a threshold value of 10
  Blue line – PDF of
  Shaded sky-blue area - the probability of observing Time to MA greater than 10
  Grey – Average and 95 % CI of

## Slide 6
Plot 3
Plot 4
Likelihood of Time to MA exceeding 10 minutes plus epsilon.
The probability of Time to MA exceeding 10 minutes in CFFH earlier decrease compared to CRTH.

## Slide 7
Plot 5
Plot 6
Similar this trend is illustrating the inverse of Cumulative Distribution functions of Time to MA accompanied by a table of percentiles.
For CRTH the probability of the 0.5 percentile is twice as large as in the CFFH assay.

## Slide 8: Distributions by Subgroups
The subsequent three slides display the comparison of distributions between CFFH and CRTH assays across the following groups:
Coagulation status: Normal, Hypo, Hyper
Timepoint: 001,002,003
Indication: Liver transplantation surgery, On-pump surgery, Not on-pump surgery
Each plot includes
raw data jittered,
box plots of data within each subgroup (Q1 and Q3)
median and mean represented by brown and grey colours, respectively.

## Slide 12
Fixed effects | Estimate | Std. error | t value | P-value
Intercept | 9.770 | 0.842 | 11.607 | <0.001
INDICATION Not On-Pump Surgery | 2.284 | 0.436 | 5.234 | <0.001
INDICATION On-Pump Surgery | 3.438 | 0.378 | 9.094 | <0.001
TIMEPOINT 002 | 1.339 | 0.179 | 7.483 | <0.001
TIMEPOINT 003 | 0.948 | 0.169 | 5.615 | <0.001
COAG STAT Hypo | -0.240 | 0.834 | -0.288 | 0.774
COAG STAT Normal | -0.099 | 0.828 | -0.120 | 0.905
Fixed effects | Estimate | Std. error | t value | P-value
Intercept | 19.762 | 0.795 | 24.869 | <0.001
INDICATION Not On-Pump Surgery | -1.940 | 0.453 | -4.284 | <0.001
INDICATION On-Pump Surgery | -1.557 | 0.392 | -3.975 | <0.001
TIMEPOINT 002 | 0.934 | 0.172 | 5.430 | <0.001
TIMEPOINT 003 | 1.926 | 0.162 | 11.868 | <0.001
COAG STAT Hypo | -0.197 | 0.779 | -0.252 | 0.801
COAG STAT Normal | 0.153 | 0.777 | 0.196 | 0.844
Model using CFFH data
Model using CRTH data
Multivariable analysis using Mixed Effect Models

## Slide 13: Additional Conclusions
For CRTH assay average Time to MA (≈20) is significantly higher than CFFH Time to MA (≈12).
  Possible explanation: Higher MA time
In both the CFFH and CRTH assays, Coagulation Status was determined to be statistically insignificant in predicting Time to MA.
Compared to the baseline at Timepoint 001, Timepoint 2 and 3 have statistically significantly higher Time to MA for both CFFH and CRTH.
Compared to baseline liver transplantation when comparing it to other indications, two distinct images emerge.
  In the CFFH assay, indications On-pump surgery, Not on-pump surgery have higher Time to MA
  Conversely, in the CRTH assay, indications On-pump surgery, Not on-pump surgery have lower Time to MA